## Slide 1: Uncertainty: the biochemistry perspective
Alison Bransfield
Bon Secours Health System

## Slide 2: Sources of variation

## Slide 3: Assessment of uncertainty for potassium
Within-run precision
Between-run precision
Calibrator
Coverage factor

## Slide 4: What is the true value??
Mean of an infinite number of measurements
No way to know the true value
  EQA
  ‘Spiked’ samples
  Reference method

## Slide 5: Within-run precision
Ensure that the instrument has been properly calibrated
Perform the test a minimum of 20 times
Calculate the mean and SD
Calculate the standard uncertainty (u) of the SD:
  u = SD/√n

## Slide 6: Between-run precision
Perform the test over a minimum of 15 runs or days
CLSI EP-12-A recommends two levels, three times a day for five days
As for within-run, calculate mean, SD and u

## Slide 7: Example:
Within-run:
  Mean = 3.56 mmol/L
  n = 30
  SD = 0.08 mmol/L
  u = 0.0146 mmol/L
Between-run:
  Mean = 3.56 mmol/L
  n = 30
  SD = 0.12 mmol/L
  u = 0.0219 mmol/L

## Slide 8: Calibration
Where the uncertainty of the calibration is known this can be incorporated
Abbott provide a guide for the Architect analyser
  Clinical Chemistry, Immunoassay Traceability, Uncertainty of Measurement. Abbott Architect 3rd edition Oct 2017

## Slide 9: Traceability of calibrator

## Slide 10: Combined uncertainty
uc = √((uwr)2 = (ubr)2 + (ucal)2)
uc = √((0.0146)2 + (0.0219)2 + (0.033)2)
= 0.0422 mmol/L

## Slide 11: Expanded uncertainty (U)
Use a coverage factor of 2 to calculate the 95% confidence interval
Thus the expanded uncertainty of a potassium measurement at 3.56 mmol/L is 0.0422 mmol/L
The ‘true’ result lies between 3.5178 and 3.6022 mmol/L

## Slide 12: Reporting uncertainty
Not practical in a typical biochemistry laboratory report
Available for users if requested – never has been…
…or has it?
Requesters will query the validity of a result or the change in a result

## Slide 13: Reference change value
RCV = 2.77 x √(CVA2 + CVI2)
  CVA = analytical variation
  CVI = within subject biological variation

## Slide 14: Thinking about uncertainty
Lab has a responsibility to be aware of the uncertainty at decision or critical levels

## Slide 15: Thinking about uncertainty
Think of other contributors to uncertainty such as dilutions

## Slide 16: Thinking about uncertainty
Tests which are calculated from other test values combine the uncertainties of all the components

## Slide 17: Thinking about uncertainty
24 hour urine collections
  Uncertainty of result
  Uncertainty of measurement of volume
  Uncertainty of collection

## Slide 18: Thinking about uncertainty
The end user is unaware of the uncertainty associated with a result
The laboratory has a duty to consider uncertainty
Performance of a method around clinical decision levels must be well defined

## Slide 19: Any questions?